- Marijuana Certified Examiner Course Equipment

Below is a list of equipment that meets the specifications for marijuana school equipment.  These products are not endorsed by the GBI-DOFS nor have most of these items been used by the Georgia Bureau of Investigation-Division of Forensic Sciences (GBI-DOFS).  This list is only a guideline and any similar product is acceptable.  Prices listed are approximate.   (effective:  Jan 2013)

- Microscopes:  Stereomicroscope, binocular eyepieces, (top) incident light, 10-40x magnification

- Fisher Standard 2000 Series    ($290)

- Fisher Student Stereo Microscope S197673 ($325)

- Fisher Scientific Stereomaster  Digital Zoom Microscopes (from $2,100)

- Edmund Scientific High Power Stereomicroscope   ($200)

- Edmund Scientific 20x/40x     ($350)

- Edmund Scientific –Boreal Digital Standard Zoom Stereomicroscope ($1,000)

- Veho Discovery VMS-004D - 400x USB Microscope with stand  ($99)

- Home Science Tools Stereomicroscope 20x/40x ($299)

- Balances:

Small Capacity:  Electronic top-loading balances, readability at least 0.1 gram, capacity up to 5000 g.

- AND Model EK3000i ($525)

- Denver Instrument Peak Series  ($200 - $500)

- Denver XP Series ($500)

- Mettler Toledo BD Series, Entry Level Series, and Portable Series  ($650 - $1600)

- Ohaus Adventurer  ($800 - $1000)

- Ohaus Scout Pro  ($200 - $450)

- Large Capacity:  Electronic top-loading, capacity greater than 5000 g (~ 11 lbs)

- Ohaus ES Bench Scales ($365 - $510)

- AND FG 30 Bench Scale ($630)

- Ohaus Champ General Purpose Bench Scale  ($510 - $1200)

- Denver Instrument XI-Series Industrial Scale ($850 - $925)

- Ohaus Champ II Bench Scales ($645 - $925)

- Balance Certification Guidelines:  Balance certification by a metrology company or other recognized vendor must be obtained prior to class.*

The balance certificate obtained must include the following information:

- Report Information:

- Model number

- Serial number

- Date

- Agency

- Calibration expiration date

- Company name

- Technician name

- Contact information

- Calibration Information:

1)  Class of weights used during calibration: ASTM Class 4 (or equivalent) or better for small capacity balances.  ASTM Class 6 or NIST 105-1 1990 class F or better for large capacity balances.)

2)  Value of weight used (ex—200.0g)

3)  Weight traceability information (NIST, National Institute of Standards and Technology)

4)  Balance readings “before” service/calibration and balance readings “after” service/calibration

* In order to operate the balance in accordance with DOFS protocol, the balance certification will be renewed annually.

Vendors:  Use certification guidelines in selection process.

- Rite-Weight, Inc.	 770-476-8500	 Toll Free 800-486-0000 	www.rite-weight.com

- Scale Systems,  Macon: 478-746-7623,  Atlanta 770-921-1988 http://www.bdindustrial.com/company/scale_systems.php

- Wiggins Scale Company 	 404-346-7090	 www.wigginsscale.com

- Balance Tech, Inc., 	 678-467-5636	http://www.balancetechinc.com/index.htm

- Department of Agriculture (contact: Jason Hitchcock) 404-656-3719

- Check yellow pages for vendor near you.

Calibration-Check Weights:  Must be obtained prior to working any cases.  They are to be used to check accuracy of balance before and after sampling cases.  Requirements for small capacity balances: NIST traceable, ASTM Class 4 or better; 1 gram weight (low end) and the largest weight the balance can weigh (high end = capacity of balance).  Requirements for large capacity balances: NIST 105-1 1990 Class F or better; 1.00 lb weight (low end) and the largest weight the balance can weigh or the largest value for which the balance will be used (high end).  Search vendors listed below or consult with Calibration Vendor.   Provided that calibration checks weights are properly maintained (i.e. handled with gloved hands, tissue, or plastic tweezers; stored in a safe location/container specifically for the storage of the check weights; not dropped or handled with bare hands or metal tweezers), they are not required to be verified annually, but can be verified at the agency’s discretion. GBI-DOFS recommends verification of the calibration check weights every three years or more frequently if significant changes occur.

- Equipment Vendors:  (for PPE, microscopes, balances, weigh boats, weigh paper, etc.)

- Rite-Weight, Inc (for balances and weights):  770-476-8500	 www.rite-weight.com

Edmund Scientific; www.scientificsonline.com   1-800-728-6999

Fisher Scientific; www.fishersci.com   1-800-766-7000

Cole Parmer; www.coleparmer.com   1-800-323-4340

VWR Scientific Products; www.vwrsp.com   1-800-932-5000

Scales Online; www.scalesonline.com   866-856-6100

Home Science Tools; www.homesciencetools.com  1-800-860-6272

- Test Kits:

Effective 07/2008, all students must bring 30 (3 boxes) of Duquenois-Levine and 30 (3 boxes) of Fast Blue B (KN Reagent) test kits.

Test kits must consist of the POUCH type DUQUENOIS LEVINE TEST and the POUCH type FAST BLUE B TEST/KN REAGENT.

Order/Part numbers are as follows:

- ODV, Inc. (NarcoPouch, NarcoTest); www.odvinc.com  1-800-852-0300

- Duquenois Levine                     NacroPouch #908        10/box

- Fast Blue B (KN Reagent)        NarcoPouch #909        10/box

- Sirchie; www.sirchie.com  1-800-356-7311

- Duquenois Levine                     NARK2005                 10/box

- Fast Blue B (KN Reagent)        NARK20020               10/box

Dana Safety Supply; 1-800-822-8722 (Carries Sirchie products)

- Duquenois Levine                     NARK2005                 10/box

- Fast Blue B (KN Reagent)        NARK20020               10/box

GT Distributors; 1-800-241-8950 (carries all brands)

- ODV

Duquenois Levine                     NacroPouch #908        10/box

Fast Blue B (KN Reagent)        NarcoPouch #909        10/box

- NIK

Test E (Duquenois Levine)        10/box

- SIRCHIE (do not regularly stock but can order and will be shipped directly from Sirchie)

- Lynn Peavey Co; http://www.lynnpeavey.com/  1-800-255-6499

- QuickCheck Pouch types

Marijuana DL QuickCheck Pouch (10/box) order number 10120

Marijuana KN QuickCheck Pouch (10/box) order number 10121

- NIK

Test E (Duquenois Levine) 	10/box

- ODV

Duquenois Levine                     NacroPouch #908        10/box            	order number 06908

Fast Blue B (KN Reagent)        NarcoPouch #909        10/box             	order number 06909